# DAWN & miniDAWN Flow Cell Cleaning

# Objective:

This course will familiarize the student with disassembly, cleaning, and assembly of a DAWN or miniDAWN flow cell.

# Introduction:

## Prevention

- Use only miscible fluids when flushing:  Toluene  100% Alcohol  Water  Aqueous buffers

- Filtration of mobile phase

  - HPLC grade solvents and 0.1 um pre column filter

  - Syringe w/ 0.02 um filter

- Prevent bacterial growth – we use 200 ppm NaN3 (absorbs at 214 nm)

- Bypass the detector when changing columns

- In SEC mode, keep the flow running 24/7, re-circulate mobile phase which helps to polish the column.  Hint:  The Wyatt solvent recycler, the Orbit, can help here.

- Use a column with little column shedding (e.g., WTC SEC columns for protein work)

- Use the Comet on a daily basis – runs about 2 hours and then shuts itself off.

- Rinse out the inlet tubing before putting a clean flow cell back in

  - (PBS in lines, toluene to calibrate – oops: re-clean the flow cell)

## Symptoms of a dirty flow cell

- Bright spots

- Increased baseline voltage and noise

| Ideal baseline and noise levels: - HELEOS II: - Noise: ~30-50 μV (mostly from photodiode, solvent independent) - Baseline: Aqueous ~ 0.01 V - Baseline: THF ~ 0.04 V - Baseline: Toluene ~ 0.16 V - HELEOS I and TREOS: - Noise: ~20-30 μV - Baseline: Aqueous ~ 0.02 V - Baseline: THF ~ 0.08 V - Baseline: Toluene ~ 0.25 V - EOS and Tristar - Noise: ~2-3 mV - Baseline: Aqueous ~ 0.2 V - Baseline: THF ~ 0.8 V - Baseline: Toluene ~ 2.5 V | baseline noise |
| --- | --- |

- Baseline wobble

- LS Overlay plot in ASTRA  System  Methods  Light Scattering  Diagnostics  detector overlay (a plot of all light scattering signals overlaid, low and high angles are typically distorted first)

All detectors overlay:

clean flow cell

Angular dependence of LS signal present for large molecules

Flat baselines:

clean flow cell

Imperfect detector overlay:

Dirty flow cell

Wavy and noisy baselines:

Dirty flow cell

## Cleaning in situ to dislodge particles

- Try a solvent with a different polarity – aqueous to organic solvent (e.g. IPA, EtOH, MeOH)

- Strong acid – 20% nitric acid, even overnight

- Weak base – NaOH pH 10 (0.1 mM!) or less (pH 12 or greater can etch the glass)

- What dissolves your sample?

    - Proteins – protease mix or enzymatic contact lens cleaner

- Comet

## Removing the flow cell from the instrument

- MiniDAWN – loosen the bolts and lift the whole assembly out

- DAWN – keep the unions with the flow cell and insert clean plugs.

- Option:  Send it back to us to clean it for $350 / included with Service Contract Annual Inspection

Flow cell parts

- Flow cell =  $3,000, no spares included; Flow cell windows = $ 300 each (2 spares included with new instrument).

- Windows have Anti-Reflective coating on the large optical surface, which can be removed by 20% nitric acid. Use nitric acid only when the flow cell is assembled!

- Small optical surface is in contact with the mobile phase

Cleaning

- Alconox (pH 9-10) or LiquiNox (pH 8-9) at 1% - follow directions on the detergent.

- Gloved finger, rub with the detergent, rinse with distilled water, rinse with IPA, blow-dry with filtered nitrogen gas.  Alternative to N2 is Dust-Off but beware of potential residue: spray to the side, not straight on.

- Sonicating – IPA okay.  Do NOT use Acids! (Removes the antireflective coating!)

## Reassemble

Remember to flush the inlet line first!

# Cell Cleaning Details

## Tools and equipment needed:

- Ball drivers--- 1.5 mm, 2.5 mm and 4 mm (for DAWN DSP and some EOS cells)

- Aperture Installation Tool---WTC p/n 119033  (for window retainer removal on current HELEOS/EOS/miniDAWN)

- Tweezers---WTC p/n 9011

- Jeweler’s Loupes----4x, 5x,  available from Edmund Scientific

- 260 mL Nalgene squeeze bottles with stream and droplet nozzles, for isopropyl alcohol (IPA)

- WTC cell cleaning kit---WTC p/n 900001

- Miniature nylon bristle brush 1/16”from Tanis Inc.---Tanis p/n 05162 @ www.tanisinc.com

- HPLC pump or syringe pump

- Compressed nitrogen system or Dust-Off dusting gas

- Oral-B dental floss---WTC p/n P8430

## Cleaning fluids:

- Isopropyl alcohol, histological grade available from VWR

- Alconox

- Deionized water

## Flow Cell Assembly Parts List:

| Item | P/N | Description |
| --- | --- | --- |
| 1 | S5002-3004 | M3 screw |
| 2 | 200694 | Manifold, out |
| 3 | P6504-20067075 | Flow cell O-ring (2) 6mm outside diameter |
| 4 | 200609 | Backing ring (2) |
| 5 | 212095 | Flow cell |
| 6 | 200690 | Manifold, in |
| 7 | P6504-20047075 | Window O-ring (2) 4mm outside diameter |
| 8 | 116007-0093 | Flow cell window (2) |
| 9 | 212073 | Window retainer (2) |
| 10 | S5002-3030 | M3 screw |
| 11 | 211048 | Bottom flow cell retainer |
| 12 | S6501 | Bottom retainer O-ring cord |
| 13 | S5002-2006 | M2 screw (2) |

| Flow Cell Disassembly Procedure: |  |  |  |
| --- | --- | --- | --- |
| - The first step in cleaning the cell is to remove the cell from the instrument. If you have a HELEOS, TREOS or EOS, disconnect the flow and unscrew the fittings at the two 0.010” Valco unions, leaving the unions attached to the two 3” tubes installed on the cell. Now remove the cell from the read head. |  |  |  |
| - If your instrument has been in use for some time, the windows will have become sealed to the 4mm o-ring inside the manifold. Using force to remove the windows may damage them unnecessarily. There are two ways you can loosen the windows from the o-rings. For the first method, schedule your cell cleaning so that you can leave your cell overnight filled with IPA and with the window retainers loosened. This overnight soak should loosen the window/o-ring seal enough that the windows simply fall out the next day. |  |  |  |
| - The second method will allow you to loosen your windows in about 15 minutes using a syringe pump. Use a 10 mL syringe and a Luer fitting to flow IPA into the outlet tube of the cell. Attach a waste line to the 0.005” inlet tube and start flowing IPA @ 1 mL/min. The 0.005” tube acts as a flow restriction and will build a small amount of pressure in the cell. Loosen one window retainer and you should see IPA begin to flow out around the retainer. Loosen and tighten the retainer several times. This compresses and relaxes the o-ring and helps loosen the manifold/o-ring/window seal. Tighten the first retainer and repeat this procedure for the other window. The pressure inside the cell forces IPA past the window/o-ring seal and allows easy window removal. At WTC we have a dedicated IPA HPLC pump that we use for this purpose. |  |  |  |
| - Now that we have loosened our windows, let’s begin the cell disassembly. Our work area can be nothing more than a clean, hard tabletop or lab bench. Wipe the area with soap and water or alcohol and dry thoroughly. |  |  |  |
| - Next, assemble the tools we will need and lay them out at the top of the table. These include the 1.5 and 2.5 mm ball drivers, the tweezers and the window retainer removal tool. |  |  |  |
| - Because we want to keep the center of the work area as clean as possible, begin cell disassembly off to the side or better yet, on a separate table. Start by blowing the IPA out of the cell using the nitrogen or Dust-Off gas being careful not to contaminate the work area. Next, use the 1.5 mm ball driver to remove the bottom cell retainer. |  |  |  |
| - Now use the 2.5 mm ball driver to remove both manifold screws. Gently separate the manifolds from the glass flow cell and carefully set the glass cell aside. It can happen that the glass cell/o-ring seal is also very tight. If that is the case, you will want to immerse the entire cell in a bath of IPA and let it soak until the glass cell can be easily released from the 6mm o-rings in the manifold. |  |  |  |
| - Still working off to the side, remove each window retainer (using the window removal tool) and use the Dust-Off to gently blow into the fluid connection opening while holding your finger over the 6mm o-ring opening. Use your thumb to prevent the window from flying out if it suddenly releases. Stubborn cases may require overnight soaking to release the window. |  |  |  |
| - The 4mm o-ring may come out with the window but it usually remains inside the manifold. To remove it, use the 1.5mm ball driver to gently pry up from the center from all sides. |  |  |  |
| - The 6mm o-rings and the 9mm back-up rings are left in place unless they are damaged or have a build-up of material between them. The 9mm backup ring can be damaged during removal and can cause leaks if it becomes deformed and is reinstalled. |  |  |  |
| Manifold Cleaning Procedure: |  |  |  |
| - We now want to clean the manifolds. Cells used with aqueous mobile phase systems often show a small amount of white salt build-up on the 6mm o-rings in the area above the bore. Take each manifold to a sink and rinse with water. At WTC we have a de-ionized water spigot at a lab sink. Now rinse with Alconox and rub the entire manifold with your gloved fingers. Take the 1/16” brush and rub the face of the 6mm o-ring. Now run the brush through the bore and use the tip to clean the threads for the window retainer. Finally, carefully run the tip of the brush into the fluid connection hole to clean the channel and spin the brush to clean the threads. Rinse thoroughly with water. Those labs without a source of running de-ionized water may simply use HPLC water in a squeeze bottle. Now rinse the manifold thoroughly with IPA from a squeeze bottle, flushing the water from all openings and surfaces. The squeeze bottle of clean IPA is a very important component of this cell cleaning procedure. We always rinse it twice before finally filling it and using it. |  |  |  |
| - Complete the cleaning of the manifold by blowing the IPA off with nitrogen or Dust-Off until manifold is completely dry. Use the 5x loupe to inspect the o-rings for damage, particles or residual deposits. The o-rings should have a uniformly smooth, flat black finish when clean. Set the dry manifold in the center of your clean area and repeat for the other manifold. |  |  |  |
| - The 9mm back up ring may have come out during the cleaning process; or if you took it out, now is the time to rinse it off, dry it and replace it in the manifold. It may fit loosely or it may be a tight fit. To insert them, simply lay them in place and use the tip of the tweezers to push into the manifold. If they are deformed or damaged, replace them with spares from the shipping kit you received with your instrument. |  |  |  |
| - Now take the two 4mm o-rings to the sink, slide them onto the 1/16” brush, wet with Alconox and spin them up and down the brush with your gloved fingers. Rinse with de-ionized water, remove them from the brush, grip both with the tweezers and rinse with de-ionized water again. Now rinse with IPA and blow dry. Again, use the 5x loupe to inspect the o-rings. At your work area, drop an o-ring into each manifold making sure it’s laying flat on the bottom of the window recess. |  |  |  |
| Flow Cell Window Cleaning Procedure: |  |  |  |
| - The next step is to clean the windows. a) The simplest method of cleaning is use the tweezers to hold the window by the stem and then wipe it on a lens tissue that has been wrapped around your finger and moistened with HPLC ethanol. b) For the more complete method, the stem of the window is held by the tweezers, soaked with Alconox, and rubbed between gloved fingers. After which additional Alconox is flushed to rinse, followed by water, then ethanol and blow-dry. Alconox was developed specifically for cleaning lab glassware and does a very good job cleaning deposits on windows. The smaller optical surface may need more than one cleaning attempt to remove all deposits. Rub gently to prevent damage to the anti-reflection coating on the larger optical surface. |  |  |  |
| - Flush the window thoroughly with de-ionized water and then rinse with IPA. We want to spray with a fairly strong jet of IPA, hitting both optical surfaces a couple times. We also want to hold the tweezers upright so that all the IPA drains down and away from the window. |  |  |  |
| - Blow-dry the window with nitrogen or Dust-Off. The best technique to use is to start the gas flowing and then, with the tweezers held almost horizontally, smoothly move the window into the gas stream. Hold the window steady in the gas flow for a couple seconds. |  |  |  |
| - While holding the window by the stem and looking toward a ceiling light, use the 5x loupe to examine the window for any particles or deposits on the optical surfaces. The ceiling adjacent to the light will be relatively dark so use this as a background for viewing the window. Against this dark background, any particles illuminated by the ceiling light will stand out as bright specks on the window surfaces. Windows should not have any cracks or chips but one or two very small particles are acceptable. |  |  |  |

| Flow Cell Window and Manifold Re-Assembly Procedure: |  |  |  |
| --- | --- | --- | --- |
| - When you are satisfied that your window is clean, set it down on the table with the large optical surface pointed to the right (if you are right-handed). Approach the window from the right side with the tweezers and grasp the window with the tweezer tips about half way down the thick portion of the window. |  |  |  |
| - Pick up a manifold and, holding it almost vertically, insert the window through the 4mm o-ring and gently release. Pull the tweezers out cleanly. The window will fit easily through the hole in the o-ring. If you grasp the window more than half way down the thick portion of the window, it will be very difficult to release the window because the tweezer tips will become jammed against the threads in the manifold. |  |  |  |
| - Now examine the installed window for particles the same way you viewed the window earlier. If all looks ok, you can now screw in the window retainer. To clean the retainers, hold both retainers with the tweezers and rinse them with IPA and blow-dry them. Set them down in your clean area and take the top one and start threading it into the manifold. Use the window retainer tool to hand tighten the window retainer until it stops, by design you cannot over-tighten the retainer. Once the stop has been reached, the retainer has bottomed out and formed a leak-proof seal between the window, o-ring and manifold. Make one last inspection of windows for particles. Repeat for the manifold. |  |  |  |
| - With both manifolds assembled, place them on the tabletop with the larger one nearest you and with the 6mm o-rings facing you. Fold seven or eight sheets of lens tissue twice in half the long way and then once the short way and lay them in the clean area of the tabletop. |  |  |  |
| Flow Cell Glass Cleaning Procedure: |  |  |  |
| - Next we’ll clean the flow cell and re-assemble it. Take the cell to the sink or wherever you are able to rinse with de-ionized water. Flush the cell with Alconox and gently roll/rub it with a gloved hand. Grasp the flow cell from the flat top and bottom surfaces to avoid fouling the optical surfaces. Run the fuzzy end of a strand of Oral-B Super Floss dental floss up from the bottom and flush the cell with Alconox so that the floss is soaked. Pull the floss back and forth three or four times and then pull it back out the bottom of the cell. Rinse thoroughly with de-ionized water as you pull the floss back-and-forth a couple more times. Then rinse thoroughly with clean IPA from the squeeze bottle. Direct a strong jet down the bore of the cell for several seconds. Rinsing the bore with a jet of clean IPA is probably the most important part of the entire procedure. Make sure the entire surface of the cell is thoroughly rinsed and finish with one more jet of IPA down the bore. |  |  |  |
| - Immediately blow-dry the cell. Again, start the flow of gas and bring the entrance of the cell bore into the flow stream. After a few seconds, move the stream away from the cell bore and try to blow off the upper and lower cell surfaces. Do not at any time stop the flow of gas while drying cell. Under these conditions the drying procedure takes about five or six seconds. It is ok if there is a little fluid remaining on the cell; this will be wiped away in the next step. At WTC we dispense our chromatographic grade nitrogen with a filter-tipped nozzle at a pressure of 40 psi. |  |  |  |
| - You should still be holding the cell with your left thumb on the step in the cell. Now take one of the folded lens tissues and drip two or three drops of IPA on the fold. Be careful at all times during this wiping of surfaces not to touch the sealing surfaces of the glass cell with lens tissue as this may leave particles or fibers that can cause leaks or get inside the bore. The tip of your finger should be near the fold. Starting in the middle of the optical surface of the cell, wipe to one end of the cell, back to the other end, once more to the other end and then back to the middle and cleanly off. At this point, you can take a clean lens tissue and wipe the top and bottom surfaces of the cell. Now, with a clean tissue, grasp the side of the cell you just wiped and flip the cell over so that the step is now on the right side and facing away from you. Repeat the wiping procedure for the remaining optical surface. |  |  |  |
| Flow Cell Complete Re-assembly Procedure: |  |  |  |
| - The cell is now ready to place on the manifold. If you have properly oriented your manifolds on the tabletop and successfully flipped your cell over during the wiping procedure you may now simply set it down on the manifolds with the step ready to engage the positioning pins on the manifold. This is why we started the procedure earlier with our left thumb on the glass step. Once you have the cell approximately positioned, turn the assembly around. Fold a lens tissue in half again and use it to grasp the cell to move the step against the pins. Insert the long manifold screw. Now bring the other manifold into position. With your thumb and two middle fingers, squeeze the two manifolds together while pushing the cell step against the positioning pins with your fore finger. Now use the 2.5mm ball driver to screw the long manifold screw almost tight. Now install and tighten the second manifold screw. Finally, go back and tighten the long manifold screw. Pick the assembly up and use the 4x loupe to sight along the glass step to be certain that the step is contacting the positioning pins. |  |  |  |
| - The final cleaning step is one last wipe of the optical surfaces. Fold a lens tissue in half three times the long way and twice the short way. Put two or three drops of IPA on the fold. With a smooth, deliberate motion, use the corner of the tissue to wipe the seam of the optical surfaces at the very ends between glass and manifold. Now take a fresh, folded tissue moistened with IPA and wipe the length of the optical surface, being careful not let your finger touch the glass as you finish. Wipe both sides in this manner. You may have to repeat once or twice to achieve a smear-free, particle-free surface. |  |  |  |
| - With the cell now clean, install the bottom retainer using the 1.5mm ball driver. The newer retainers have a wavy, spring steel washer in the center and a Viton strip at the edge to ensure correct positioning of the glass cell in the manifold assembly. Be sure that the retainer is installed so that the Viton strip pushes the glass step toward the positioning pins and not away from them. Also be certain that there is minimal overhang of the retainer on the manifold as this may cause laser alignment error when the cell is reinstalled in the read head. |  |  |  |
| - Your cell is now clean. Cap it with clean plugs and reinstall in read head. Be absolutely certain that you have flushed the inlet and outlet tubes with clean mobile phase before you reattach them to your clean cell. We recommend rinsing and blow drying all fluid fittings before connecting, especially those connections made to the flow cell. |  |  |  |

In-Situ MALS Cleaning Protocol (Aqueous)

- Before disassembling the flow cell or sending the instrument/flow cell back to Wyatt Technology, you can follow this quick cleaning protocol.

- Disconnect MALLS detector from the downstream detectors and route the outlet tubing to an appropriate waste container.

- Activate the COMET if available.

- Flush the LC-MALLS system with filtered water to remove any remaining buffer.

- After flushing with filtered room temperature water, continue to flush the system with ~500 ml of hot water (~40) followed by a 10% SDS solution.

- For tough cleaning, use a detergent solution to flush system containing a 1% Hellmanex solution (Hellma Online) or 1% Thomas Scientific Contrad 70 (Cat#2901-c15).

- Follow up detergent flush with an additional 500 ml of hot water.

- In addition to the detergent cleaning protocol, it had been shown that some sticky proteins require an enzyme cleaning step to fully remove adsorbed contaminants.

- Disconnect MALLS detector from the downstream detectors and route the outlet tubing to an appropriate waste container.

- Activate the COMET if available.

- Connect a standard syringe pump with a 20 ml syringe and flush system with room temp filtered water.

- Introduce a pre-mixed enzyme contact lens cleaning solution such as Bausch & Lomb Renu (any brand enzyme contact lens cleaning solution should work).

- Let sit overnight and flush with filtered water and/or a detergent such as those listed above.

- If you would like to use Nitric Acid as has been traditionally recommended by Wyatt Technology:

# Flush the flow cell with dilute nitric acid** (20% or any concentration from 0.01 to 0.1 Molar).

# Leave the nitric acid in the flow cell for a period of time from 15 minutes to overnight.

**Be aware that nitric acid is very caustic and should be handled with extreme caution.

# Protease Cocktail

Some users have found that a simple protease “cocktail” rinse is effective in removal of protein deposits from glass flow cell surfaces.  You might be able to use this rinsing treatment rather than disassembling the flow cell:

For 3 ml of protease cleaning solution.  All enzymes are sequencing grade preparations from either Boerhringer Manheim or Roche.

Ingredients:

- Trypsin, modified - 25 μg, lyophilized 
 - Chymotrypsin - 25 μg, lyophilized
 - Pepsin - 25 μg, lyophilized
   (You might be able to get away with just pepsin alone, as it's so non-specific.)

Procedure:

- Reconstitute each with 1 mL of PBS (25 mM Na phosphate / 150 mM NaCl, pH 7.25

- Mix the three solutions and vortex, load syringe fitted with 0.02 um filter for LS detector

- Flush detector with 20mL pure water, then infuse ~ 1 mL of cocktail via syringe pump

- Stop flow, turn on Comet (if you have one) and leave it for a few hours or overnight

- The following morning, remove syringe, flush with fresh 20 mL of water, then mobile phase.